## AC outline
The Isolette device provides safe and effective incubation care for the infant, under the management of a qualified care-giver (nurse) when operated in an environment conforming to specified environmental assumptions | Risk Management
System Scope and Purpose | Define boundary and purpose to determine scope of risk management effort
Clinical Intended Use
System Boundary | System boundary indicates what the manufacture has control over and what they don't. Also identifies the inputs, outputs, and inportant state that must be constrained. Identifies interactions with environment that must be consider for possible harmful actions
Environmental (Context) Entities | Identifies entities in the enviroment about which we will consider in risk assessment, including entities that we will make assumptions above
Environmental (Context) Assumptions
Operational Concepts and Clinical Functions | Indicates shared stakeholder understanding of how the system is intended to work in its environment
Risk Analysis
Hazardous Situations Identified | What situations (considering both system state/actions and environment) can possibly lead to harm
Initial Risk Assessment of Hazardous Situations | Ranking of hazardous situations according to risk
Risk Controls for Initial Hazardous Situations | Measures taken to reduce those risk
Risk Assessment and Residual Risk
Risk Control Functional Requirements
Risk Control Functional Requirements
Design / System Architecture Specification
System Architecture is specified
External interfaces are specified at a level of detail that supports analysis of safety/security and assessment of implementation conformance to architecture
((type of external interface))
For each physical interface
For each interoperability interface..
For each form of information flow across system boundary (monitored and controlled variables), data types, expected ranges, and units are specified
Functions implemented in hardware / software
System Architecture is specified at a level of detailed sufficient for supporting hazard analysis supporting residual risk assessment
System Architecture is specified at a level of detailed sufficient for support Common Cause Analysis

## AC Component Template
Component AC Template
| Development Phase | Task Responsibility | Quality Compliance Responsibility | Evidence Artifact | Attestation | Traceability | Model-based Development Impact | Relationship to NRC / IFE Certifacation
The Deployed Strongly-Compliant Component Implementation Satisfies it Requirements
Component Specification - Interface and Requirements
[Interface Specification] The interface of the component is specified in terms of input and output ports | Comp, Design | Comp. Developer | Self/Audit | [Model] Comp. Interface |  |  | Development quality improvement - interface is precisely specified in a abstract form (higher-level than code), which is machine readable and can be the subject of further analysis (e.g., type-checking, unit checking, enforcement of data type process rules (linking to a project data dictionary) and promotes shared understanding. This enables inconsistencies in design to be discovered with machine help early in the development process (the design phase)
[Port Specification] Each port in the interface specification is fully specified (wrt type, units, properties, types trace to "data dictionary" for the "unit")) | Comp, Design |  | Self/Audit/Resolute | [Model] Port declarations in model |  |  | (as above)
[Component Functional Requirements] The functional requirements of the component are stated in terms of the component's interface and the publicly disclosed component local state | Comp. Design | Comp. Developer | Self/Audit | [Req. Document] Subsection for component |  |  | Requirements / Specification quality improvement - natural language requirements are stated in terms of a precisely specified interface
[Component Functional Requirements Formalization] Interface requirements are formally specified as model contracts | Comp. Design | Comp. Developer | Self/Audit | [Model] GUMBO Contracts for component |  | GUMBO clauses trace to requirement(s). Each requirement traces to one or more clauses. | Verification quality improvement - natural language requirements are transformed to an unambigous format which can be (a) analyzed at the model level for semantic inconsistencies (b) inference of desired system or integration properties. This is also part of the process (the foundation) for ensuring machine-verified conformance of application code to requirements (via both automated-testing and verification)
[Component Timing Requirements] The timing requirements for the component are specified | Comp. Design | Comp. Developer | Self/Audit | [Model] Timing properties for component |  |  | Verification quality improvement - natural language requirements are transformed to an unambigous format which can be (a) analyzed at the model level for semantic inconsistencies (b) inference of desired system or integration properties (real-time constraints, latency analysis). This includes support for scheduling analysis and schedule generation.
Component Implementation -- Application Code
[Application code communicates exclusively through APIs] The application code of the component receives/sends its external inputs and outputs exclusively through APIs representing corresponding to the component's declared interface (Rationale: (1) The requirements stated in terms of the declared interface and state can be appropriately interpreted as observable constraints on source code behavior. (2) Supports information flow conformance) | Comp. Impl. | HAMR (Code generation and build). | HAMR Platform / Audit | [Infra. Code] Link to "bridge" code; APIs and implementation for component; [App. Code] Link to "includes/import" of APIs. | API code (Bridge) is up-to-date and not modified. | Bi-directional traceability between component interface declarations and API methods | Traceability / Audit / Development Management improvement -- Auto-code generation guarantees that the boundaries and interfaces of application code conform to the model-declared abstractions. If the model changes, code level APIs also change, and the developer is forced to reconcile to ensure that models and code are in sync. Each concept in the code interface is traceable to the model-level interface. The fidelity of information flow, hazard analysis, testing, etc are all improved because we enforce the exact means by which a component interacts with its context and the mechanisms and nature of the interactions are accurately captured in the model.
[Application code conforms to computational model] When dispatched, the component computes a function from its inputs and local state to outputs and updated local state, where the means of dispatching and inputs visible at dispatch adhere to the AADL standards description of thread dispatch modes (Note: perhaps part of this should be relocated to infrastructure code. What this does support is the sound interpretation of the requirements as constraints on input/output ports and local state. Perhaps more important here is reflecting the entry point methods as the root of component behaviors, with the contracts indicating what it reads and writes) | Comp. Impl. | HAMR (Code generation and build). |  | [App Code]: link to compute entry points, [Infra. Code]: link to invocation of compute entry point | Infrastructure code (dispatching) is up-to-date and not modified | Bi-directional traceability between component declaration (and dispatch protocol) and code entry point
Component Verification (Application code conforms to Requirements)
[Application code behavior satisfies functional requirements] The function computed by the component's application code satisfies the component's stated functional requirements | Comp. Verification | Distributed across HAMR Code and Contract Generation, Developer (e.g., completeness of tests), Contract Verification (Verus tool), Testing infrastructure (propTest)
Testing
[Application code behavior satisfies functional requirements - testing] Testing demonstrates that the function computed by the component's application code satisfies the component's stated functional requirements as captured in GUMBO contracts | Comp. Verification
[GUMBOX executable contracts are correct]. Oracles used for component testing (GUMBOX executable contracts) correctly implement the semantics of the GUMBO contracts | Tool infrastructure Development | HAMR Code and Executable Contract Generation. | HAMR Platform/Audit | Link to component GUMBOX contract | GUMBOX contracts are up to date and have not been modified. | Bi-directional traceability GUMBO contract clauses and GUMBOX functions implementing the clause
[Appropriate Coverage Criteria Defined] Coverage criteria for both executable specification and code is defined and justified to be appropriate for the component (considering criticality, complexity, etc.) | Comp. Verification (Planning) | Component Developer / Process Engineer | Self/Audit | Link to text description of criteria and justification (Verification Plan)
[Coverage Criteria Measurement] An appropriate method (e.g., an automated coverage tool) is applied to get an accurate coverage measurement according to the specified coverage criteria | Comp. Verification (Planning) / Tool Infrastructure | Component Developer / Process Engineer | Self/Audit | Link to text description of justification (This is essentially a placeholder for a large tool qualification step) (Verification Plan)

## Boundary
Boundary and Context Entities
Physical Interactions | Type (input/output) | Context interactions | Physical Interpretation
Current Temperature | Monitored | Air (+ infant indirectly) | Current air temperature inside Isolette
Heating Element | Controlled | Air (+ infant indirectly) | Heating element whose actuation increases the air temperature within the Isolette
Operator Interactions | Type (input/output) | Context interactions | Interpretation
Upper Desired Temperature | Monitored | Operator | Upper bound of target temperature range; Input (keyed) by Operator
Lower Desired Temperature | Monitored | Operator | Lower bound of target temperature range; Input (keyed) by Operator
Upper Alarm Temperature | Monitored | Operator | Upper bound of alarm temperature range; Input (keyed) by Operator
Lower Alarm Temperature | Monitored | Operator | Lower bound of alarm temperature range; Input (keyed) by Operator
Display Temperature | Controlled | Operator | Value of system's understand of the current temperature within the Isolette; visual output to operator
Regulator Status | Controlled | Operator | Status (reflects current mode) of Regulate Function
Monitor Status | Controlled | Operator | Status (reflects current mode) of Monitor Function
Context entities
Nurse
Air (in Isolette)
Infant (indirect, but listed since the primary safety objective pertains to the infant)
Boundary and Environment Diagrams

## Goals
Goals / Assumptions
Intended Use (simplfied)
The Isolette device provides safe and effective incubation care for the infant, under the management of a qualified care-giver (nurse) when operated in an environment conforming to specified environmental assumptions
Functional Goals
Top-Level Goals
G1 | The air temperature within the Isolette should be maintained at a level that is safe for the infant and effective for its neonatal development
G2 | The nurse shall be notified when the air within the Isolette reaches an unsafe temperature
Top-Level Environmental / Deployment Assumptions
D1 | The Isolette is operated by a nurse with appropriate qualifications and the ability to manage the Isolette's settings and alarms
D2 | The Nurse's workflow (as directed by the healthcare delivery organization) guarantees sufficient availability to manage the Isolette
D3 | The Isolette is operated in a physical and technical environment appropriate for its clinical function

## TL Sys Reqs
System Level Safety Functions
Regulate Heat Function | Comments | Anticipated Verification Method and Evidence | Impact of Model-based Development
Design (Function inputs/outputs) | Some aspect of design is necessary to characterize the inputs, outputs, and state of the function, in order to be able to properly constrain the function, and to be able to write testable requirements. Here we reference the boundary elements (monitored and controlled) relevant for the function and its inputs and expose necessary internal state maintained by the function. | Verified by inspection of the design documents and implementation. We must confirm that, e.g., there are no other inputs to the function. This enables us to include that we design our tests and verification for the function, we are systematically exercising and covering its behavior. | Model-based development and the proposed HAMR refinement framework (together will seL4 partitioning) will guarantee that we can precisely specify input and output characteristics (types, units, behavioral constraints). In implementations generated from these, we can produce traceability information that helps us locate exactly the source code and execution elements that realize the input and output -- guaranteeing that we know exactly where to probe and observe the system state to see if the behavior conforms to our requirements. Moreover, seL4 protections guarantee that we only have those precisely declared inputs and outputs.
Inputs
current temperature
upper desired temperature
lower desired temperature
Outputs
heating element state
Internal State (read; written)
The characterization of the heat regulate function shall be modal according to the following modes (implemented in stored state): - Init - Regulate components are initializing, valid temperature is acquired - Normal - no internal failures, all inputs valid, - Failed - time out on initialization, internal failure, or input failure has occured
Requirements
Init Mode Requirements
REG-Heat-Control-1: If the Regulator Subsystem mode is Init, then the heat element is Off
Normal Mode Requirements
Intent: The air within the Isolette shall be maintained at a safe temperature, according to inputs for the desired temperature range (set points) specified by the operator
REG-Heat-Control-2: If the Regulator Subsystem mode is Normal and the current_temp is less than the lower desired temp, then the heat element is On
REG-Heat-Control-3: If the Regulator Subsystem mode is Normal and the current_temp is greater than the upper desired temp, then the heat element is Off
REG-Heat-Control-4: If the Regulator Subsystem mode is Normal and the current_temp is less than or equal to upper desired temp, and the current_temp is greater than or equal to the lower desired temp, then the heat element state is unchanged
Failed Mode Requirements
REG-Heat-Control-5: If the Regulator Subsystem mode is Init, then the heat element is Off
Mode Transition Requirements
TBD. Specify conditions under which the function updates its mode
Timing Requirements
Changes to the set point inputs shall be reflected in the function's output within an interval of XX seconds
Changes to the Isolette air temperature shall be reflected in the function's behavior within an interval of 1 seconds
Changes to the mode shall be reflected in the functions output within an interval of XX seconds
Monitor Function
(omitted for now)
Independence Requirements
The Regulate Heat function and the Monitor function shall be independent in the sense that a failure in any constituent of one function's implementation shall not give rise to a failure of other (needs rewording)

## Comp Reqs
System Level Safety Functions
Regulate Heat Function
Design (Function inputs/outputs)